|  |  |
| --- | --- |
| TITLE: | **Urine Dipstick Quality Control Record Sheet** |
| ACTIVE DATE: | 01/09/2024 |

***\*\**\*NEW URINE DIPSTICK CONTROL SOLUTIONS\*\*\***

* Please find enclosed fresh QC samples for the urine dipstick reader.
* **These are to be opened / started on the 1st of the month.**
* **These QC samples are stable for 1 month:**
* The samples **must** be dated when opened.
* The samples **must** be kept in a fridge.
* Overleaf, you will find a Quality Control record sheet for use this month.
* **Only one QC sample needs to be run per day.**
* Alternate the testing between NEG & POS as on the record sheet.
* Please discard the previous month’s in-use QC samples on the 1st.
* The used QC can be discarded as for clinical waste.
* **In case of a QC failure, please contact POCT on (01225 82) 6044.**

**Urine Dipstick Quality Control Record Sheet**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Location: |  | Serial No: |  | Month/Year |  |
| NEG QC Lot No: |  | Strip Lot No: |  |  |  |
| POS QC Lot No: |  | Strip Expiry: |  |  |  |

***NOTE: Only one urine QC is run each day, alternating between Negative & Positive.***

|  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **Date** | **QC** | **Nit** | **Leu** | **Ket** | **Glu** | **Bld** | **Prt** | **pH** | **SG** | **Cleaned** | **Signature** |
| 1 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 2 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 3 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 4 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 5 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 6 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 7 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 8 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 9 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 10 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 11 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 12 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 13 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 14 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 15 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 16 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 17 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 18 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 19 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 20 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 21 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 22 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 23 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 24 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 25 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 26 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 27 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 28 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 29 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| 30 | *POS* |  |  |  |  |  |  |  |  |  |  |
| 31 | *NEG* |  |  |  |  |  |  |  |  |  |  |
| **QC Ranges** | **Nit** | **Leu** | **Ket** | **Glu** | **Bld** | **Prt** | **pH** | **SG** | **Ranges** |
| **Negative** | NEG | NEG | NEG | NEG | NEG | NEG | 5 - 6 | 1.020 -1.030 | **Negative** |
| **Positive** | POS | 15 -500 | 0.5 -30 | 8.3 -56 | 10 -250 | 0.3 -5.0 | 7.0 -9.0 | 1.005 -1.015 | **Positive** |

|  |  |
| --- | --- |
| **At the end of every month, please scan this sheet and email it to POCT using:****ruh-tr.biochempoc@nhs.net****For queries, phone ext. 6044** | **Notes / Comments:** |