| TITLE: | Urine Dipstick Quality Control Record Sheet |
| --- | --- |
| ACTIVE DATE: | 01/09/2024 |

***NEW URINE DIPSTICK CONTROL SOLUTIONS***

- Please find enclosed fresh QC samples for the urine dipstick reader.

- These are to be opened / started on the 1st of the month.

- These QC samples are stable for 1 month:

- The samples must be dated when opened.

- The samples must be kept in a fridge.

- Overleaf, you will find a Quality Control record sheet for use this month.

- Only one QC sample needs to be run per day.

- Alternate the testing between NEG & POS as on the record sheet.

- Please discard the previous month’s in-use QC samples on the 1st.

- The used QC can be discarded as for clinical waste.

- In case of a QC failure, please contact POCT on (01225 82) 6044.

Urine Dipstick Quality Control Record Sheet

| Location: |  | Serial No: |  | Month/Year |  |
| --- | --- | --- | --- | --- | --- |
| NEG QC Lot No: |  | Strip Lot No: |  |  |  |
| POS QC Lot No: |  | Strip Expiry: |  |  |  |

NOTE: Only one urine QC is run each day, alternating between Negative & Positive.

| Date | QC | Nit | Leu | Ket | Glu | Bld | Prt | pH | SG | Cleaned | Signature |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | NEG |  |  |  |  |  |  |  |  |  |  |
| 2 | POS |  |  |  |  |  |  |  |  |  |  |
| 3 | NEG |  |  |  |  |  |  |  |  |  |  |
| 4 | POS |  |  |  |  |  |  |  |  |  |  |
| 5 | NEG |  |  |  |  |  |  |  |  |  |  |
| 6 | POS |  |  |  |  |  |  |  |  |  |  |
| 7 | NEG |  |  |  |  |  |  |  |  |  |  |
| 8 | POS |  |  |  |  |  |  |  |  |  |  |
| 9 | NEG |  |  |  |  |  |  |  |  |  |  |
| 10 | POS |  |  |  |  |  |  |  |  |  |  |
| 11 | NEG |  |  |  |  |  |  |  |  |  |  |
| 12 | POS |  |  |  |  |  |  |  |  |  |  |
| 13 | NEG |  |  |  |  |  |  |  |  |  |  |
| 14 | POS |  |  |  |  |  |  |  |  |  |  |
| 15 | NEG |  |  |  |  |  |  |  |  |  |  |
| 16 | POS |  |  |  |  |  |  |  |  |  |  |
| 17 | NEG |  |  |  |  |  |  |  |  |  |  |
| 18 | POS |  |  |  |  |  |  |  |  |  |  |
| 19 | NEG |  |  |  |  |  |  |  |  |  |  |
| 20 | POS |  |  |  |  |  |  |  |  |  |  |
| 21 | NEG |  |  |  |  |  |  |  |  |  |  |
| 22 | POS |  |  |  |  |  |  |  |  |  |  |
| 23 | NEG |  |  |  |  |  |  |  |  |  |  |
| 24 | POS |  |  |  |  |  |  |  |  |  |  |
| 25 | NEG |  |  |  |  |  |  |  |  |  |  |
| 26 | POS |  |  |  |  |  |  |  |  |  |  |
| 27 | NEG |  |  |  |  |  |  |  |  |  |  |
| 28 | POS |  |  |  |  |  |  |  |  |  |  |
| 29 | NEG |  |  |  |  |  |  |  |  |  |  |
| 30 | POS |  |  |  |  |  |  |  |  |  |  |
| 31 | NEG |  |  |  |  |  |  |  |  |  |  |
| QC Ranges |  | Nit | Leu | Ket | Glu | Bld | Prt | pH | SG | Ranges |
| Negative |  | NEG | NEG | NEG | NEG | NEG | NEG | 5 - 6 | 1.020 -1.030 | Negative |
| Positive |  | POS | 15 -500 | 0.5 -30 | 8.3 -56 | 10 -250 | 0.3 -5.0 | 7.0 -9.0 | 1.005 -1.015 | Positive |

| At the end of every month, please scan this sheet and email it to POCT using: ruh-tr.biochempoc@nhs.net For queries, phone ext. 6044 | Notes / Comments: |
| --- | --- |