| Department/Location/Project: PAW theatres & ITU use of TEG6 analyser and TEG Manager. | SOP Document Reference Number: SOP/POCT/7/3 |
| --- | --- |
| Risk Assessor(s): Nicola Hodges & Jonny Price | Highest Risk Rating Identified*:15 |
| Date of assessment:16/06/2021 | Informed QM of any Risk Score >9 (NH): |

* Any identified risk which has a rating >9 must be communicated with the Quality Manager

| Description of risk | Existing control/ safe System of work | Initial Risk Rating (S X L= RR) |  |  | What further action is required | Responsible person and target date for completion | Final Risk Rating (S X L= RR) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| The process of acquiring a venous blood sample could cause the following: -Needle stick injuries -Infection -Blood spillages due to puncture wounds Affecting the clinical staff and patient. | - Training for venepuncture - Needle stick injury policy - Immunizations - PPE - Sharps Disposal policy - Procedure for spillages of body fluids - Pathology Health and safety policy - Training for treatment of wounds. | 2 | 3 | 6 |  |  | 2 | 3 | 6 |
| Analysis of samples or external quality control can pose a risk of infection to ODPs, anaethetic nurses and anaesthetists. | - PPE - Disposal into contaminated waste bins | 1 | 3 | 3 |  |  | 1 | 3 | 3 |
| Inadequately filled clotting sample (citrate) can produce inaccurate results This could result in an unnecessary treatment being given to a patient (e.g. blood products) or withholding of treatment (blood products) when actually required for a patient. | - Online ‘College of TEG’ training - User guide explicit - Laboratory results are still taken and reviewed. These will be utilized to further confirm/refute the TEG 6 results. Face to face TEG6 training - SOP & competency assessment in place - Access control allowing only trained & competent staff to use the analyser. - Traceability of analysis. | 1 | 5 | 5 |  |  | 1 | 5 | 5 |
| Misinterpretation of results that fall above or below the action limits. A potential risk that abnormal result may not get acted on appropriately and coagulopathy may be missed. Blood products may be given to a patient when not required. | - Online ‘College of TEG’ training - Clear prompts are set up on the TEG 6 to guide action in the case of abnormal results. - How to act on an abnormal result is covered in ‘College of TEG’ training. - Laboratory results are still taken and reviewed. These will be utilized to further confirm/refute the TEG 6 results. - Face to face training - Lab access to TEG 6 output (via TEG Manager) to cross reference/validate results - Access control so only competent staff can use analyser - Link/super user in place to support trouble shooting of abnormal results. | 1 | 5 | 5 |  |  | 1 | 5 | 5 |
| Patient ID not being entered into meter before blood is analysed can lead to: - Incomplete audit trail for that patient so trends cannot be viewed reliably. - One ID number used for more than one patient. These can lead to the wrong results being assigned and acted upon for a patient. Leading to unnecessary treatment (e.g. blood products) or withholding of treatment (blood products) from a patient when actually required. | - TEG 6 requires patient MRN for test - TEG log asks for username and patient ID per test - Process described in ‘College of TEG’ and user guide | 2 | 5 | 10 | - TEG 6 link into Trust IT systems. As such access and patient details will be auto-populated via bar codes - Audit of TEG user list - INTAUD/POCT/227 (Aug 2021) - Vertical Audit to be scheduled for 6 months for both analysers (ITU & PAW) audit reference: INTAUD/POCT/207 – To be completed by POCT with support from departments. | Paperless project will hopefully see connectivity of TEG results with the EPR - 2022 August 2021 July 2021 | 1 | 5 | 5 |
| Generic log on/password for analyser Difficult to audit individual staff use of analyser. This can lead to untrained staff using the analyser possibly producing incorrect results. This may effect which products are given or not to the patient. No ownership or responsibility of results produced. | - Username and password established by Haemonetics - Online ‘College of TEG’ training - User guide explicit Individual issue of username and passwords for TEG 6 post training face to face or College of TEG and competency assessment. - Maintain a record of trained users. - Put in place reassessment of user competency. | 1 | 5 | 5 | Removal of generic password | September 1st 2021 post audit INTAUD/POCT/227 | 1 | 5 | 5 |
| Transcription of results into patient notes The result could be written incorrectly from the analyser to the notes, leading to unnecessary treatment (e.g. blood products) or withholding of treatment (blood products) from a patient when actually required | - During training users are taught to write the correct result in the patient record and sign this transcript - Laboratory results are also reviewed if unexpected POC results are produced. - The printout is checked against the transcription by a second person. | 3 | 5 | 15 | - Link into Trust Cerner EPR to allow real time digital recording or results. As such, hand-written notes will become redundant | Trusts Paperless project will hopefully see connectivity of TEG results with the EPR - 2022 | 1 | 5 | 5 |
| Damaged or faulty analyser could produce incorrect results. Leading to unnecessary treatment (e.g. blood products) or withholding of treatment (blood products) for a patient when actually required. | - TEG 6 has a robust system for self-testing with every test run. - TEG 6 will have an IQC performed weekly (as per SOP) and not be used if QC fails. - If TEG 6 is down and results cannot be obtained the procedure previously in place will be followed using laboratory results. - Registration with EQA scheme ukneqas assures the accuracy of the results produced on a quarterly basis. - Arrangements of use of PAW theatre analyser or Paw theatre to use ITU analyser as a contingency/backup - Maintenance and service contract set up with manufacturer | 1 | 2 | 2 |  |  | 1 | 2 | 2 |
| Electrical Hazard risk of electrocution to user (staff). This could be caused by a sample / fluid spillage onto TEG 6 or through faulty connection to mains electricity. | - PAT testing performed on all hospital equipment in the department. | 1 | 1 | 1 | - Confirm PAT test current | Local – ensure this is included with whole department | 1 | 1 | 1 |
| There is a risk to patients who have complicated clotting pathologies (like that seen in COVID19 infection) of a delay in receipt of blood products or incorrect products given if staff are unable to access the analyser. | - College of TEG training available to staff required to use analyser. - All staff required to use analyser must be trained. - Link trainers in each department with administrator access to TEG manager for adding users | 3 | 2 | 6 | - Review of user access time period to ensure 2 yearly reassessment - 6 monthly audits to be scheduled to get an updated list from haemonetics of completed competencies and update onto TEG manager. | NH - INTAUD/POCT/227 August 2021 | 3 | 1 | 3 |
| TEG Manager there is a risk of server downtime possibly causing: - Logins not to work - Loss of data - No remote access to results | - TEG manager sits on a VM server, if the trust server goes down all systems will go down. If the VM goes down it will restore itself automatically. - Results can still be viewed on the analysers themselves. - The analysers will retain the settings and data from the previous synchronisation. | 2 | 1 | 2 |  |  | 2 | 1 | 2 |
| If the network goes down TEG6 analysers will not be able to talk to TEG manager, a risk of results not being stored electronically and not accessible through TEG manager | - The network going down will affect all IT aspects of the trust. The analyser will still be able to be used in the same. The data will be stored on the analyser and resent when the network is re connected. | 3 | 1 | 3 |  |  | 3 | 1 | 3 |

- Risk assessment matrix

Acceptable Risk

Risk is tolerable as long as it is well managed and controlled.  In addition to identified hazards, all incidents claims and complaints will be risk assessed according to the following process and investigated according to the severity or the consequence and likelihood of (re)occurrence.

All Risk Assessments within the Trust will identify:

- The hazards within the Task/ area being assessed inherent in the work undertaken

- who and how many people would be affected

- how often specific events are likely to happen (may be based on frequency of previous occurrence):

- how severe the effect or consequence would be

- how controllable the hazards are.

Acceptable risk will be determined using the following traffic light system:

Severity/consequence

Given the (in) adequacy of the control measures, how serious the consequences are likely to be for the group, patient or Trust if the risk does occur (using the matrix).

|  | Consequence score (severity levels) and examples of descriptors |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | 1 | 2 | 3 | 4 | 5 |
| Domains | Negligible | Minor | Moderate | Major | Catastrophic |
| Impact on the safety of patients, staff or public (physical/ psychological harm) | Minimal injury requiring no/minimal intervention or treatment. No time off work | Minor injury or illness, requiring minor intervention Requiring time off work for ≤3 days Increase in length of hospital stay by 1-3 days | Moderate injury requiring professional intervention Requiring time off work for 4-14 days Increase in length of hospital stay by 4-15 days RIDDOR/agency reportable incident An event which impacts on a small number of patients | Major injury leading to long-term incapacity/ disability Requiring time off work for >14 days Increase in length of hospital stay by >15 days Mismanagement of patient care with long-term effects | Incident leading to death Multiple permanent injuries or irreversible health effects An event which impacts on a large number of patients |
| Quality/complaints/ audit | Peripheral element of treatment or service suboptimal Informal complaint/inquiry | Overall treatment or service suboptimal Formal complaint (stage 1) Local resolution Single failure to meet internal standards Minor implications for patient safety if unresolved Reduced performance rating if unresolved | Treatment or service has significantly reduced effectiveness Formal complaint (stage 2) complaint Local resolution (with potential to go to independent review) Repeated failure to meet internal standards Major patient safety implications if findings are not acted on | Non-compliance with national standards with significant risk to patients if unresolved Multiple complaints/ independent review Low performance rating Critical report | Totally unacceptable level or quality of treatment/service Gross failure of patient safety if findings not acted on Inquest/ombudsman inquiry Gross failure to meet national standards |
| Human resources/ organisational development/ staffing/ competence | Short-term low staffing level that temporarily reduces service quality (< 1 day) | Low staffing level that reduces the service quality | Late delivery of key objective/ service due to lack of staff Unsafe staffing level or competence (>1 day) Low staff morale Poor staff attendance for mandatory/key training | Uncertain delivery of key objective/service due to lack of staff Unsafe staffing level or competence (>5 days) Loss of key staff Very low staff morale No staff attending mandatory/ key training | Non-delivery of key objective/service due to lack of staff Ongoing unsafe staffing levels or competence Loss of several key staff No staff attending mandatory training /key training on an ongoing basis |
| Statutory duty/ inspections | No or minimal impact or breech of guidance/ statutory duty | Breach of statutory legislation Reduced performance rating if unresolved | Single breech in statutory duty Challenging external recommendations/ improvement notice | Enforcement action Multiple breeches in statutory duty Improvement notices Low performance rating Critical report | Multiple breeches in statutory duty Prosecution Complete systems change required Zero performance rating Severely critical report |
| Adverse publicity/ reputation | Rumours Potential for public concern | Local media coverage – short-term reduction in public confidence Elements of public expectation not being met | Local media coverage – long-term reduction in public confidence | National media coverage with <3 days service well below reasonable public expectation | National media coverage with >3 days service well below reasonable public expectation. MP concerned (questions in the House) Total loss of public confidence |
| Business objectives/ projects | Insignificant cost increase/ schedule slippage | <5 per cent over project budget Schedule slippage | 5–10 per cent over project budget Schedule slippage | 10–25 per cent over project budget Schedule slippage Key objectives not met | Incident leading >25 per cent over project budget Schedule slippage Key objectives not met |
| Finance including claims | Small loss Risk of claim remote | Loss of 0.1–0.25 per cent of budget Claim less than £10,000 | Loss of 0.25–0.5 per cent of budget Claim(s) between £10,000 and £100,000 | Uncertain delivery of key objective/Loss of 0.5–1.0 per cent of budget Claim(s) between £100,000 and £1 million Purchasers failing to pay on time | Non-delivery of key objective/ Loss of >1 per cent of budget Failure to meet specification/ slippage Loss of contract / payment by results Claim(s) >£1 million |
| Service/business interruption Environmental impact | Loss/interruption of >1 hour Minimal or no impact on the environment | Loss/interruption of >8 hours Minor impact on environment | Loss/interruption of >1 day Moderate impact on environment | Loss/interruption of >1 week Major impact on environment | Permanent loss of service or facility Catastrophic impact on environment |

Likelihood

Given the (in) adequacy of the control measures for each risk, decide how likely the risk is to happen according to the following guide.  Scores range from 1 for rare to 5 for very likely.

| Score | Descriptor | Description |
| --- | --- | --- |
| 1 | Rare | Extremely unlikely to happen/recur – may occur only in exceptional circumstances – has never happened before and don’t think it will happen (again) |
| 2 | Unlikely | Unlikely to occur/reoccur but possible. Rarely occurred before, less than once per year. Could happen at some time |
| 3 | Possible | May occur/reoccur. But not definitely. Happened before but only occasionally - once or twice a year |
| 4 | Likely | Will probably occur/reoccur. Has happened before but not regularly – several times a month. Will occur at some time. |
| 5 | Very Likely | Continuous exposure to risk. Has happened before regularly and frequently – is expected to happen in most circumstances. Occurs on a daily basis |

Risk Score is determined by Severity x Likelihood

|  | Consequence |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Likelihood | 1 Insignificant | 2 Minor | 3 Moderate | 4 Major | 5 Catastrophic |
| 5 – Almost certain | 5 | 10 | 15 | 20 | 25 |
| 4 - Likely | 4 | 8 | 12 | 16 | 20 |
| 3 – Possible | 3 | 6 | 9 | 12 | 15 |
| 2 – Unlikely | 2 | 4 | 6 | 8 | 10 |
| 1 - Rare | 1 | 2 | 3 | 4 | 5 |