## EQA failures
Yearly Quality summary of POCT 2021.
[9 empty] | Not includng Dec |  | Not including Dec
FPP meters - failures over 3 distributions |  | Blood gas | Failures out of 66 samples each | Pregnancy | Failures out of 18 samples each | ID Now | one sample per month | Urinalysis | Failures out of 33 samples each | Hemocue | Failures out of 22 samples each (10 for transfusion) | Coaguchek | Failures out of 10 samples each | TEG | Failures out of 3 samples | HbA1C - Affinion | Failures out of 33 samples | I-STAT | Failure out of 6 samples
27 non returns | 1 poor performance | Lab | 0 | ED | 2 | ED | 0 | ED | 1 | Theatre 11 | 4 | Gastro Ward | 0 | PAW | 1 (interpretation only) | Children's OP | 0 | EPAC | 1 (non return)
cardiac 2 | ITU | ITU 2 | 5 (pO2) | CHW | 5 | ONCOP | 0 | CHW OP | 0 | PAW theatre | 5 | Eye Ward | 1 | ITU | 0
haygarth 3 |  | BBC | 1 (lac) | Charlotte | 0 | CHW | 0 | CHW | 0 | Main PACU | 3 | X-Ray | 0
pain clinic 1 |  | NICU | 0 | Dermatology | 0 | MAU | 0 | CHW extra | 2 | PAW theatre | 3 | Pro II Coag nurses Nurses | 0
pulteney 1 |  | ED | 2 (glu & OxHb), 1 (carboxyHb), 1 (pCO2) | GUM | 0 | OPAU | 0 | Urology 1 | 1 | Main PACU 2 | 3
wb 3 |  | MAU | 5 (glu), 1 (K) | Gynae OP | 0 | SAU | 0 | Urology 2 | 2 | Day surgery | 3
robin smith 2 |  | Resp | 7 (oxyHb), 6 (metHb), 2 (carboxyHb & tHb) | Haygarth | 0 | BBC | 0 | Parry | 0 | Transfusion 1 | 0
ED 3 |  | ITU 3 | 2 (pO2) | ITU | 0 | Chip mat | 0 | GUM | 3 | Transfusion 2 | 0
mau 1 |  | ITU 1 | 4(pO2) & 1 (glu) | Nuclear med | 0 | Frome mat | 0
eye ward 2 |  | SAU | 3 (coox), 1 (ca), 1 (glu), 1 (pO2) | SAU | 1
chw 2 |  | Cath lab | 2 (pO2),5 (lac), 6 (Glu), 1(H+,Na,pCO2) | SSSU | 0
combe 1 |  | ITU 4 | 6 (pO2), 1(Hb) | X-ray | 1
mary 1 |  | CHW | 6 (glu), 4 (po2), 2 (tHb), 1 (OXHb), 1(metHb)
asu 1
water house 1
sssu 1
pierce 1
main theatres 1

## CAPA
Number | Details | Fault Category | Severity | Raised Against | Owner | Raised Date | Closed Date
EQA/454 | Urine EQA distribution W123 saw one non return (charlotte Ward) and 2 incorrect results returned (SAU & Children's ward). | Procedures\Procedures not followed | Minor | POCT | Hodges, Nicola | 2021-01-19 | 2021-04-08
EQA/457 | Blood gas EQA distribution 303 (November). ITU (3 &4) - each had only one level run on it. ie both sample one's were run on ITU (4) and both samples twos' run on ITU(3). They were not reported in WEQAS due to this so do not have any performance data. Looking back at the reports for this distribution ITU (4) performed well on sample 1 and ITU (3) performed well for sample 2. | Pre-examination Error | Minor | POCT | Hodges, Nicola | 2021-01-21 | 2021-04-14
VAL/POCT/30 | Validation of two Abbott ID Now PCR analysers for COVID 19 testing on ED. | Validation | Minor | POCT | Hodges, Nicola | 2021-02-02 | 2021-04-21
VAL/POCT/31 | Re-validation of the Omni-s blood gas analyser. | Validation | Minor |  | Hodges, Nicola | 2021-03-12 | 2021-03-18
EQA/472 | Theatre 1 has failed to return a result for their hemocue meter for EQA distribution 2103Hb. The ward indicated that the meter had been lost and was why they were unable to analyse the EQA sample on the meter. | Equipment | Minor | POCT | Hodges, Nicola | 2021-03-24 | 2021-05-14
EXTAUD/UKAS/530 | This is a UKAS finding for the Omni-S blood gas analyser located in the laboratory "Maintenance logs for the Cobas 221 daily (QMS/FM/POCT/77/1) and monthly maintenance (QMS/FM/POCT/78/!) are not always completed for daily maintenance schedules.". finding reference number: 219840-03-01-E01902-001 |  |  | POCT | Hodges, Nicola | 2021-03-24 | 2021-06-14
EQA/466 | we had a poor performance result for the Cath lab blood gas EQA for glucose sample 1. we reported 2.2 when the result from the analyser was 22. | Procedures | Minor | POCT | Hodges, Nicola | 2021-04-01 | 2021-04-01
EQA/467 | External quality assurance samples were delivered to the children's ward on the 21/01/2021 (distribution no. 304) and signed for by a member of staff. The samples were processed the following day and the report returned to the laboratory 27/01/2021. The results were entered by laboratory staff but there were no printouts to reliably second check transcription so the computer screen was used but not all parameters are visible this way. On return of the performance report there were several parameters that did not score well pO2, glucose, total haemoglobin, methaemoglobin, and oxyhaemoglobin. These poor results may be the result of inadequate mixing or preparation of the samples. | Procedures\Procedures not followed | Minor | POCT | Hodges, Nicola | 2021-04-06 | 2021-04-06
EQA/468 | External quality assurance samples (WEQAS) were delivered to the Respiratory ward on the 21/01/2021 (distribution no. 304) and signed for by a member of staff. The samples were processed the following day and the report returned to the laboratory 27/01/2021. The results were entered by laboratory staff into the WEQAS website. On return of the performance report there were several parameters that did not score well pO2, total haemoglobin, methaemoglobin, carboxyhaemoglobin and oxyhaemoglobin. These poor results may be the result of inadequate mixing or preparation of the samples. | Procedures\Procedures not followed | Minor | POCT | Hodges, Nicola | 2021-04-06 | 2021-04-06
EQA/469 | External quality assurance samples were delivered to the children's ward on the 18/02/2021 (distribution no. 305) and signed for by a member of staff. The samples were processed on the 20/02/2021 and the report returned to the laboratory 22/02/2021. The results were entered by laboratory staff but there were no printouts to reliably second check transcription. On return of the performance report there were 2 parameters that did not score well pO2 & total haemoglobin. These poor results may be the result of inadequate mixing or preparation of the samples. Also the returned results sheet says the analyser was calibrating but this at worse would only have affected one sample. | Procedures\Procedures not followed | Minor | POCT | Hodges, Nicola | 2021-04-06 | 2021-05-14
FSN/25 | FAN 915-416 for ABL90 Flex and ABL90 Flex plus - Risk of incorrect time and potentially biased results. This does relate to our blood gas analysers on the ITU wards in particular 092R0246N031 as it does not have an on board battery. JA performed the checks confirming there is either a fitted battery or an external battery for the analyser. | Supplier\Field safety notice | Minor |  | Hodges, Nicola | 2021-04-09 | 2021-04-09
EQA/474 | Children's Ward reported a poor result for urine pregnancy testing EQA distribution W124. Sample 3 was reported as negative when the sample was a spike urine sample with HCG at 300IU/L which should have been a clear positive result. This is the second poor result in a row for Children's ward. | Procedures\Procedures not followed | Minor | POCT | Hodges, Nicola | 2021-04-22 | 2021-04-22
EQA/475 | Urine pregnancy EQA distribution W124: ED had a poor performance for sample 3 where they had reported a negative result when the sample was a spiked urine sample with HCG level of 300 IU/L which should have been a clear positive. | Procedures\Procedure followed (SOP incorrect) | Minor | POCT | Hodges, Nicola | 2021-04-22 | 2021-04-22
INC/HS/112 | The ED blood gas analyser was in a very dirty state on Friday there was blood over the surface, screen, floor, cupboard and sample port area. People were using it in this state and making no attempt to clean it. It was very difficult to use the analyser without standing in the blood or touching the blood on the analyser. | Health & Safety | Minor | POCT | Hodges, Nicola | 2021-05-04 | 2021-05-04
FSN/29 | This is a bulletin to remind users of the Roche b123 blood gas analyser to return security levels back to level 4 if it was changed during the pandemic. | NO FAULT - IMP/UAT/VAL (other) |  |  | Hodges, Nicola | 2021-05-10 | 2021-05-10
EQA/479 | Day surgery reported 78.0 g/L when the target value was 82.79 g/L giving a DI score of -3.33 | Process failure\Quality control\EQA failure | Other | POCT | Hodges, Nicola | 2021-05-11 | 2021-07-16
EQA/476 | THe April WEQAS distribution for glucose and ketone EQA was performed entirely by the staff on POCT. This was to identify the location of all the meters as over the past 12 months we have seen meters moving between wards frequently. During this distribution we were unable to fine 7 meters despite repeat attempts and investigations. This means these meters did not have EQA run on them this time. | Process failure\Quality control\EQA failure | Minor | POCT | Hodges, Nicola | 2021-05-13 | 2021-05-13
EQA/477 | RAU had a poor performance for distribution 306 (March) blood gas EQA. Glucose sample 1 SDI = -3.17, Sample 2 SDI = -2.54, Sample 3 SDI = -1.42 overall SDI = 2.38. The iQC fine, MSS replaced on 4th no obvious causes lactate OK but out of range for sample 3 (no result reported). Ca Sample 1 SDI = 2.05 overall acceptable SDI = 1.33. | Training | Other | POCT | Hodges, Nicola | 2021-05-17 | 2021-06-24
EQA/488 | Theatre 12 and main PACU HB2 DI>2.0 | Equipment\Records\Records not complete | Other | POCT | Hodges, Nicola | 2021-06-23 | 2021-07-16
EQA/485 | Children's ward reported poor results for April and May pO2 EQA. This has been an issue for all of 2021. Emails and datixs raised. | Procedures\Procedures not followed | Minor | POCT | Hodges, Nicola | 2021-06-25 | 2021-10-12
VAL/IT/65 | Re-validation of the Comark gateway unit SN:26140027 | NO FAULT - IMP/UAT/VAL (other) |  |  | Hodges, Nicola | 2021-06-28
EQA/486 | The external quality assurance (EQA) samples were taken to Respiratory ward on the 21/05/2021 to be analysed on their blood gas analyser by a member of staff. The results were not returned until the final date 1/6/21 and I believe the samples were all left at room temperature during this time (3 of the 6 should be kept in the fridge). The report was issued several weeks later and the performance of several parameters on the COOX module was poor. The samples are not stable outside of the fridge and is the likely cause for the poor performance. | Procedures\Procedures not followed | Other | POCT | Hodges, Nicola | 2021-06-29 | 2021-06-29
VAL/POCT/35 | Re-Validation of the TEG6 analyser on ITU to establish the MoU using 12 months of iQC data. | Validation | Minor | POCT | Hodges, Nicola | 2021-06-30 | 2021-07-02
EQA/490 | Several wards failed to return results for the July FPP EQA: Children s ward: KAAU150-D0400, KAAW068-A0077 glucose and ketone Cardiac ward meter KAAV313-A0514 Glucose Haygarth ward meter KAAV336-A0646 for glucose and ketone MAU meter KAAU013-D2164 for glucose William budd day case meters KAAV329-A0122, KAAX103-A0220 for glucose ED meter KAAV202-A0220 for glucose. | Training | Minor | POCT | Hodges, Nicola | 2021-08-03 | 2021-08-03
COMP/HOS/108 | Just wanted to say a Big Thanks to you and your Great Team on the much Support they gave us on resolving our pxp barcode issues. Great work Emily, Jan & Dan great Team work Nic you should be proud of them all they were right on it getting this situation resolved for us. Can you please say a Great Big Thank you to them all for there much needed help it has been very much Appreciated Nic :-) Maureen Avery, Technical Assistant, Neonatal unit zone D2 |  |  | POCT | Hodges, Nicola | 2021-08-18 | 2021-08-18
EQA/493 | Theatre 11 and Main PACU both returned a poor result for the hemocue EQA 2105HB. | Procedures\Procedures not followed | Minor |  | Hodges, Nicola | 2021-08-18 | 2022-01-11
AUD/POCT/99 | Follow up Wards with greater than 10% usage. ? Repeat offenders ? Training records ? Email Link trainers Checklist:Usage of Emergency Numbers on FPP meters Question No. 1.5 | Training | Minor | POCT | Hodges, Nicola | 2021-09-01 | 2022-01-12
SUPPLIES/482 | An MSS decontamination has been left for 13 days until Roche can send an engineer this has meant lactate and glucose has been unavailable over the weekend and will not be working until Tuesday. | Supplier | Other |  | Hodges, Nicola | 2021-09-06 | 2021-10-12
EQA/492 | Children wards Urine pregnancy EQA for July had one sample reported as negative when it was from a 17 week pregnant urine donor. This is the fourth failed urine pregnancy EQA in the past five distributions. | Procedures\Procedures not followed | Other | POCT | Hodges, Nicola | 2021-09-07 | 2021-09-09
EQA/494 | Urine pregnancy EQA was not returned by ED, after a phone call the samples could not be found and our spare set had already been given to another department so there was no way to send out replacement samples. | Procedures | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-10
EQA/496 | Theatre 11 and PAW hemocue EQA results were >than 2DI. | Training | Other | POCT | Hodges, Nicola | 2021-09-22 | 2021-10-26
COMP/HOS/110 | Thank you for “fixing” the ID Now analysers in the ED department. From Simon Sethi, Chief Operating Officer. |  |  | POCT | Hodges, Nicola | 2021-11-08 | 2021-11-08
UAT/POCT/10 | Acceptance for use of level 3 auto QC for the Roche b221 blood gas analyser product number: 03321185001, lot: 21400967, exp: 31/12/21. Calibrations had passed and the QC performed within the manufacturers stated SD's | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-07 | 2021-04-07
UAT/POCT/4 | Waste bottle 21403403, 31/07/2023 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-03-23 | 2021-03-23
UAT/POCT/5 | User acceptance for use of S1 wash solution pack lot: 21403201, EXP: 01/08/2022 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-03-23 | 2021-03-23
UAT/POCT/6 | Acceptance of S3 lot:21402305, exp: 29/11/2021 on NICU BGA | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-03-31 | 2021-03-31
UAT/POCT/7 | Tube holder complete (peristaltic pump tubing) product 03087603001 lot: 464020, exp: not stated. Passed acceptance criteria | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-03-31 | 2021-03-31
UAT/POCT/8 | S1 was solution product 03260917184 lot:21402804, exp: 04/07/2022 went into use on the 26/03/2021 on the laboratory blood gas analyser. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-03-31 | 2021-03-31
UAT/POCT/9 | MSS cassette for the measurement of lactate and glucose on the Roche b221 blood gas analyser product number: 03261085184, lot:21510248, exp: 30/04/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-03-31 | 2021-03-31
UAT/POCT/11 | Acceptance for MSS cassette for the Roche b221 blood gas analyser product number: 03261085184, lot: 21510347, exp: 07/05/21. Calibrations had passed and the QC performed within the manufacturers stated SD's | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-07 | 2021-04-07
UAT/POCT/12 | Acceptance for use of level 1 auto QC for the Roche b221 blood gas analyser product number: 03321185001, lot: 21400967, exp: 31/12/21. Calibrations had passed and the QC performed within the manufacturers stated SD's | CAPA Raised in Error* | Minor | POCT | Hodges, Nicola | 2021-04-07 | 2021-05-14
UAT/POCT/13 | S1 solution 03321169001, lot: 21490772, exp: 30/04/2021 from delivery 04/03/2021 acceptance into use. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-15 | 2021-04-20
UAT/POCT/14 | AutoTrol iQC level 2, product number: 03321177001, lot: 21400870, Exp: 31/12/21. Received on 01/04/2021 (5 boxes). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/15 | pCO2 electrode product code: 03111679180, lot 21504347, exp 23/04/2021 (install before). Received 9/01/2021 (2 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/16 | pH electrode product number 03111717180, lot 21505047, install before 11/06/2021. Received 24/03/2021 (4 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/17 | Tube holder complete (peristaltic pump tubing) product number 03087603001, lot 464020 no install or expire date stated. Received 07/12/2020 one of 22 units. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/18 | Calcium electrode for the Roche b221 analyser, product number 03111644180, lot 21504347, install before 23/04/2021. Received 19/01/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/19 | pO2 electrode for the Roche b221 blood gas analyser product number 03111695180, lot 21504547. Received 19/01/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/20 | Reference electrode for Roche b221 blood gas analyser product number 03111873180, lot 21504647, install before 14/05/2021. Received 24/03/2021 (8 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/21 | Potassium electrode for Roche b221 blood gas analyser product number 03111628180, lot 21504848, install before 28/05/2021. Received 24/03/2021 (s units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/22 | Sensor contact electrode for the Roche b221 blood gas analyser product code 03260909184, lot 21502347, install before 03/06/2022 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/23 | Dummy electrode for the Roche b221 blood gas analyser product number 03111849035, lot 21594847, install before 26/11/2021. Received 19/01/2021 (2 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/24 | Sodium electrode for the Roche b221 blood gas analyser product number 03111598180, lot 21504547, install before 07/05/2021. Received 19/01/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/25 | Reference contact electrode for Roche b221 blood gas analyser product number 03112071180, lot 21502247, install before 27/05/2022. Received 07/12/2020 (5 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/26 | AutoTrol iQC level 1 for the Roche b221 blood gas analyser product number 03321169001, lot 21490772, exp 30/04/2021. Received 01/04/2021 (2 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-21 | 2021-04-21
UAT/POCT/27 | THB Cuvette for the Roche b221 blood gas analyser product number 03086879, lot 200047, expiry: 15/05/2025. Received 11/11/2020 (8 units). Calibrated with calibrator lot: 21493801, epiry: 31/08/2022 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-27 | 2021-04-27
UAT/POCT/28 | Fill port for the Roche b221 blood gas analyser product number: 04847229001, lot: 324020. Received 05/11/2020 (8 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-04-27 | 2021-04-27
UAT/POCT/29 | Auto QC level one for use on the Roche blood gas analysers product number: 04015630017386, lot 21490772, exp: 30/04/21 . Was accepted into use on the 27/04/21 on the ED analyser. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-07 | 2021-05-07
UAT/POCT/30 | Auto QC level 1 for use on the Roche blood gas analyser product number: 04015630017386, lot 21490772, exp: 30/04/2021. Was accepted into use on the NICU analyser 01/05/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-07 | 2021-05-07
UAT/POCT/31 | pCO2 electrode for the Roche b221 blood gas analyser product number: 0311679180, lot number: 21505048, install before: 11/06/21. Received 24/03/21 (3 units) and was accepted into use on the 29/04/21 on the ED analyser. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-07 | 2021-05-07
UAT/POCT/32 | S3 solution for the Roche b221 blood gas analyser product number: 03260933184, lot: 21403801, exp: 14/03/2022. Received 04/03/2021 (2 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-07 | 2021-05-07
UAT/POCT/33 | pO2 electrode for the Roche b221 blood gas analyser product number: 03111695180, lot: 21505047, install before: 11/06/2021. Received 24/03/2021 (3 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-07 | 2021-05-07
UAT/POCT/34 | MSS cassette for Roche b221 blood gas analyser product number: 03261085184, lot 215010747, install before: 04/06/2021. Received 05/05/2021 (8 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-07 | 2021-05-07
UAT/POCT/35 | S2 solution pack for the Roche b221 blood gas analyser product number: 03260925184, Lot: 21404402, expiry: 25/04/2022. Received on the 04/03/2021 (5 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-07 | 2021-05-07
UAT/POCT/36 | AutoQc level 1 for the Roche b221 blood gas analyser product number 03321169001, lot: 21400778, exp: 08/22. Received 05/05/21 (14 mats). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-07 | 2021-05-07
UAT/POCT/43 | Waste bottles for the Roche b221 blood gas analyser product number 03144054001, lot 21410507, expiry: 31/01/2023. Received 05/05/2021 (10 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-13 | 2021-05-13
UAT/POCT/44 | Autotrol QC level 3 product number 03321185001, lot: 21400969, expiry: 31/05/2022. Received 05/05/2021 (4 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-13 | 2021-05-13
UAT/POCT/45 | S3 solution for the Roche b221 blood gas analyser product code: 03260933184, lot: 21403807, expiry: 14/03/2022. Received 04/03/2021 (5 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-13 | 2021-05-13
UAT/POCT/46 | pO2 electrode for the Roche b221 blood gas analyser product number 0311169518, lot 21505047, install before 11/6/21. Received 06/05/21 (5 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-18 | 2021-05-18
UAT/POCT/47 | T&D disk for the Roche b221 blood gas analyser product number: 05039428001, lot 534020, no expiry stated, received 07/12/20 (8 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-18 | 2021-05-18
UAT/POCT/48 | Peristaltic pump tubing (tube holder) for the Roche b221 blood gas analyser product cone: 03087603001, lot: 464020, no expiry stated, received 06/01/2021 (22 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-18 | 2021-05-18
UAT/POCT/49 | Sensor contact for Roche b221 blood gas analyser product number 03260909184, lot: 21503047, install before 22/07/22, received 07/12/20 (10 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-18 | 2021-05-18
UAT/POCT/50 | Reference electrode for the Roche b221 blood gas analyser product number: 03111873180, lot: 21505147, install before 18/06/2021, received 06/05/2021 (9 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-18 | 2021-05-18
UAT/POCT/51 | Calcium electrode for the Roche b221 blood gas analyser product number: 03111644180, lot 21504649, install before 14/05/2021, received 24/03/2021 (2 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-18 | 2021-05-18
UAT/POCT/52 | pH electrode for the Roche b221 blood gas analyser product 03111717180, lot 21505047, install before 11/6/21, received 06/05/2021 (6units). The NICU analyser (SN:16527) was having some issues before this electrode was replaced and after. It was not possible to evaluate this electrode due to this and it was replaced before this electrode passed QC on the analyser. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-05-18 | 2021-05-19
UAT/POCT/53 | Auto trol level 3 QC for the Roche b221 blood gas analyser product code: 03321185001, lot: 21400969, expiry: 31/05/22. Received 02/06/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-02 | 2021-06-22
UAT/POCT/54 | MSS cassette for the Roche b221 blood gas analyser product code: 03261085184, lot: 21511147, expiry: 02/07/2021. Received 02/06/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-02 | 2021-06-22
UAT/POCT/55 | Waste bottles for the roche b221 blood gas analyser product code: 03144054001, lot: 21410507, expiry: 31/01/2023. received 02/06/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-02 | 2021-06-09
UAT/POCT/56 | Auto-trol level 1 QC for the Roche b221 blood gas analyser product code: 03321169001, lot: 21400778, expiry: 31/08/2022. Received 02/06/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-02 | 2021-06-04
UAT/POCT/57 | Auto-trol level 2 QC for the Roche b221 blood gas analyser product number: 03321177001, lot: 21400873, expiry: 31/08/2022. REceived 02/06/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-02 | 2021-06-22
UAT/POCT/58 | S3 solution pack for the Roche b221 blood gas analyser product code: 03260933184, lot: 21410216, expiry: 11/07/2022. Received 02/06/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-02 | 2021-06-09
UAT/POCT/59 | S1 solution pack for the Roche b221 blood gas analyser product code: 03260917184, lot: 21410205, expiry: 09/01/2023. REceived 02/06/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-02
UAT/POCT/60 | Printer paper for the roche b221 blood gas analyser product code: 03113361180, lot: 32755, expiry: 03/02/2028. REceived 02/06/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-02
UAT/POCT/61 | S2 solution pack for the Roche b221 blood gas analyser product code: 03260925184, lot: 21410419, expiry: 25/07/2022. REceived 02/06/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-02 | 2021-06-22
UAT/POCT/62 | S3 solution for the Roche b221 blood gas analyser product number 03260933184, lot: 21410216, expiry: 11/07/2022. Received 05/05/2021 (4 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/63 | Sodium electrode for the Roche b221 blood gas analyser product code: 03111598180, lot: 21504647, expiry: 14/05/2021. Received 24/03/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/64 | Dummy electrode for the Roche b221 blood gas analyser product code: 03111849035, lot: 21594847, expiry: 26/11/2021. Received 24/03/2021 (8 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/65 | S2 solution for the Roche b221 blood gas analyser product code: 03260925184, lot: 21410314, expiry: 18/07/2022. Received 06/05/2021 (10 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/66 | Sodium electrode for the Roche b221 blood gas analyser product code: 03111598180, lot: 21505047, expiry: 11/06/2021. Received 06/05/2021 (5 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/67 | Calcium electrode for the Roche b221 blood gas analyser product code: 03111644180, lot: 21505047, expiry: 11/06/21. Received 06/05/2021 (5 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/68 | Potassium electrode for the Roche b221 blood gas analyser product code: 03111628180, lot 21504848, expiry: 28/05/2021. Received 06/05/2021 (5 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/69 | Reference contact electrode for the Roche b221 blood gas analyser product code: 03112071180, lot: 21503047, install before: 22/07/2022. Received 24/03/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/70 | Reference contact electrode for the Roche b221 blood gas analyser product code: 03112071180, lot: 21503447, install before: 19/08/2022. Received 24/03/2021 (1 unit) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/71 | T&D disk for the Roche b221 blood gas analyser product code: 05039428001, lot: 484220, expiry: NA. Received 06/01/2021 (8 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/72 | Sensor contact electrode for the Roche b221 blood gas analyser product code: 03260909184, lot: 21503947, install before 23/09/2022. Received 24/03/2021 (6 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/73 | pCO2 electrode for the Roche b221 blood gas analyser product code: 03111679180, lot: 21505049, install before: 11/06/2021. Received 06/05/2021 (5 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/74 | Fillport for the Roche b221 blood gas analyser product code: 04847229001, lot: 324020, expiry: NA. Received 07/12/2020 (8units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-03 | 2021-06-03
UAT/POCT/75 | Auto-trol QC level 2 product code: 03321177001, lot: 21400873, expiry: 03/02/2022 (4 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-04 | 2021-06-04
UAT/POCT/76 | S1 solution for the Roche b221 blood gas analyser product code: 03260917184, lot: 21404306, expiry: 17/10/2022. Received 04/02/2021 (6 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-04 | 2021-06-04
UAT/POCT/84 | Reference contact electrode for the Roche b221 blood gas analyser product number: 03112071180, lot: 21503447, install by: 19/05/2022. Received 06/05/2021 (4 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-15 | 2021-06-15
UAT/POCT/85 | pH electrode for the Roche b221 blood gas analyser product number: 031717180, Lot: 21510547, install before: 06/08/2021. Received 07/05/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-22 | 2021-06-22
UAT/POCT/86 | Waste separator for the Roche b221 blood gas analyser product number: 05783445001, lot: 324020. Received 17/11/2020 (5 units) | NO FAULT - IMP/UAT/VAL (other) | Other | POCT | Hodges, Nicola | 2021-06-22 | 2021-06-22
UAT/POCT/92 | S1 rinse solution for the Roche b221 blood gsa analyser product number: 03260917184, lot: 21404306, expiry: 17/10/2022, received 04/02/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-06-30 | 2021-06-30
UAT/POCT/105 | SD bypass nipple for the Roche b221 blood gas analyser product code: 03274969001, lot: 393819 exp: NA. accepted into use on the Respiratory analyser 07/07/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-07-28 | 2021-07-28
UAT/POCT/107 | S1 solution for the Roche b221 blood gas analyser product number: 03260917184, lot: 21404306, exp: 17/10/2022. Delivered 04/03/2021 (3 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-08-03 | 2021-08-03
UAT/POCT/108 | Docking mechanism pack for the waste bottle on the b221 roche blood gas analyser product code: 05414563001, lot: 10695463. Received 07/05/2021 2 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-08-04 | 2021-08-04
UAT/POCT/109 | Level 1 iQC for the Roche b221 blood gas analyser product number: 3110812022, lot: 21400778, exp: 02/08/2022 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-08-04 | 2021-08-04
UAT/POCT/110 | THB cuvette for the Roche b221 blood gas analyser product number: 03086879001, lot: 190207, exp: 28/06/2024. Received 03/12/2020 (8 units). | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-08-26 | 2021-08-26
UAT/POCT/118 | Waste bottles for Roche b221 blood gas analyser product number: 03144054001, lot: 21411905, expiry 30/04/2023 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/119 | Level 3 iQC for the ROche b221 blood gas analyser product number: 03321185001, Lot: 21400969, Expiry: 31/05/2022. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/120 | Level 3 iQC for the Roche b221 blood gas analyser product code: 03321177001, lot: 21400873, expiry: 31/08/2022 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/121 | Level 1 iQC for the Roche b221 blood gas analyser product code: 03321169001, lot: 21400778, expiry: 31/08/2022. Received 05/08/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/122 | MSS cassette for the Roche b221 blood gas analyser product code: 03261085184, lot: 21512247, expiry: 17/09/2021, received 05/08/2021 (7units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/123 | Ph electrode for the Roche b221 blood gas analyser product code: 03111717, lot: 21511864, Install by: 05/11/21, received 24/08/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/124 | MSS cassette for the Roche b221 blood gas analyser product code: 03261085184, lot: 21512647, expiry: 15/10/2021, received 03/09/2021 (7 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/125 | S3 solution for the Roche b221 blood gas analyser product code: 03260933184, lot: 21410903, expiry: 29/08/2022 received: 05/08/2021 (4 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/126 | Reference electrode for the Roche b221 blood gas analyser product code: 03111873, lot: 21511755, install by: 29/10/2021. Received 24/08/2021 (4 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/127 | pO2 electrode for the ROche b221 blood gas analyser product code: 03111695, lot: 21511862, Install by: 05/11/2021. Received 24/08/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/128 | pCO2 electrode for the Roche b221 blood gas analyser product code: 03111679, lot: 21511863, install by: 05/11/2021. Received 24/08/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/129 | S1 solution for the Roche b221 blood gas analyser product code: 03260917184, lot: 21411612, expiry: 17/04/2023. Received 05/08/2021 | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/130 | S2 solution for the Roche b221 blood gas analyser product code: 03260925184, lot: 21411807, expiry: 31/10/2022. Received 05/08/2021 (10 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/131 | S1 solution for the Roche b221 blood gas analyser product code: 03260917184, lot: 21404306, expiry: 17/10/2022. Received: 04/03/2021 (2 units) | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-09-09 | 2021-09-09
UAT/POCT/133 | MSS cassette for the ROche b221 blood gas analyser product code: 03261085184, lot 21513247, install by: 11/2021. Delivered 08/10/2021 7 units. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-10-13 | 2021-10-13
UAT/POCT/134 | Auto-trol iQC level 2 for the Roche b221 blood gas analyser product code: 03321177001, lot: 21400873, exp: 31/08/2022. Delivered 02/09/2021 4 units. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-10-13 | 2021-10-13
UAT/POCT/135 | S3 solution for the Roche b221 blood gas analyser product code: 03260933, lot: 21411607, exp: 10/2022. delivered 26/08/2021 one unit | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-10-13 | 2021-10-13
UAT/POCT/136 | Tube holder complete (pump tubing) for the Roche b221 blood gas analyser product number: 03087603001, lot number: BP2228E, expiry date: NA. Delivered 27/10/2021 25 units. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-11-17 | 2021-11-17
UAT/POCT/137 | Auto-trol level 3 product number 03321185001, lot number: 21400969 exp: 05/2022. Delivered 08/10/2021 4 units. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-11-17 | 2021-11-17
UAT/POCT/138 | MSS cassette product number 03261085184, lot: 21513347, install by: 03/12/21. Delivered 03/11/21 7 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-11-17 | 2021-11-17
UAT/POCT/172 | S1 solution pack for the Roche b221 blood gas analyser product number 03260917184, lot 2141311 use by 31/7/23. Delivered 03/11/2021 9 units. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-03 | 2021-12-03
UAT/POCT/176 | clot catchers for use with samples used on the Roche b221 blood gas analysers, product code: 03112012180, lot: 21267421, exp: 28/06/2024, deliv: 03/11/2021 20 units. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/177 | tHB Cuvete for the Roche b221 blood gas analyser product: 03086879001, lot: 190207, exp: 28/06/2024. Delivered: 07/1/2021 8 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/178 | Needle for Roche b221 blood gas analyser product: 03309746001, lot: 363320 no expiry. Delivered 17/11/20 2 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/179 | Waste container product 03144054001, lot: 21413434, expiry: 31/7/23. Delivered 03/11/2021 12 units. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/180 | Autotrol level 1 iQC product 03321169001, lot: 21400778, exp: 31/08/2022. delivered: 03/11/21 7 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/181 | S2 solution for the Roche b221 blood gas analyser product: 03260925184, lot: 21413112, exp: 30/01/2023. Delivered 03/11/2021 14 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/182 | S3 solution for the ROche b221 blood gas analyser product: 03260933184, lot: 21412307, exp: 05/12/22. Delivered 08/10/21 12 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/184 | S2 solution for the Roche b221 blood gas analyser product 03260925184 , lot: 21412606, exp: 12/22. Delivery 08/10/21 14 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/185 | S3 solution for the Roche B221 blood gas analyser product 03260933184, lot: 21411607, exp: 10/22. delivered: 27/08/21 1 unit | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/186 | waste container product 03144054001, lot: 21411905, exp: 04/23. delivery 08/10/21, 13 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/187 | S3 solution for the Roche b221 blood gas analyser product: 03260933184, lot: 21412707, exp: 02/01/23. Delivery 03/11/21 12 units. | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-10 | 2021-12-10
UAT/POCT/196 | Fill port for the Roche b221 blood gas analyser product: 04847229001, lot: 434120, expiry 06/01/2022. Delivery: 06/01/2021 8 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-22 | 2021-12-22
UAT/POCT/197 | Dummy Electrode for the Roche b221 blood gas analyser product: 03110192035, lot: 21500547, expiry: 28/01/2022. unsure date of delivery 2 units | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-22 | 2021-12-22
UAT/POCT/199 | S3 fluid pack for the Roche b221 blood gas analyser product code: 03260933184, lot: 21412707, exp: 02/01/23. Deliv: 27/10/21 1 unit | NO FAULT - IMP/UAT/VAL (other) | Minor | POCT | Hodges, Nicola | 2021-12-30 | 2021-12-30

## Datix
Web83542 | 2021-01-19 | Charlotte ward did not return results for Urine pregnancy testing EQA W123/ | CA/PA = EQA/454 | Fully approved
Web83754 | 27.1.21 | SAU returned incorrect result for urine pregnancy EQA W123 | CA/PA = EQA/454 | Fully approved
Web83755 | 27.1.21 | Children’s ward returned an incorrect result for urine pregnancy EQA W123 | CA/PA = EQA/454 | No investigation 22/4/21
Being approved 6/5/21
Closed 3/6/21
Web85736 | 2021-04-06 | Childrens ward returned several poor results for pO2, glucose, total haemoglobin, methaemoglobin, and oxyhaemoglobin. Blood gas EQA dist 304 | CA/PA = EQA467 | No investigation 6/5/21
2021-06-03
Closed 16/6/21
web85740 | 2021-04-06 | Respiratory ward – poor performance on EQA 304 (January 21) for pO2, tHb, COhb, oxyHb, methb | CA/PA = EQA468 | No investigation 6/5/21
Awaiting approval 3/6/21
Finally approved 14/6/21
web |  | Children’s ward gas EQA distribution 305. Sample 2 pO2 SDI = 2.88, sample 3 pO2 SDI = 2.55, overall SDI = 2.09. Return said machine calibrating. THb sample 3 SDI = -4.45 overall SDI = 1.72 | CA/PA =
Web86076 | 2021-04-20 | The EQA was taken to theatres on 29/03/2021, they were unable to analyse the samples on one of their meters (20005H-1512013015) theatre 1 and had claimed this meter was lost. This is now the second distribution in a row that this meter has been missing. | CA/PA = EQA/472 | No investigation 6/5/21
2021-06-03
2021-06-29
Web86077 | 2021-04-20 | Hemocue EQA dist 2103Hb – poor results for PACU 2 DI = -2.46. | CA/PA = EQA/471 | No investigation 6/5/21
2021-06-03
Web86088 | 2021-04-20 | Audit findings INTAUD/POCT/202 – hemocue documentation and record keeping still not in place a year on from previous audit and datix. | CA/PA =TREND/27 | No investigation 6/5/21
2021-06-03
Web86140 | 2021-04-22 | Children's Ward reported a poor result for urine pregnancy testing EQA distribution W124. Sample 3 was reported as negative when the sample was a spike urine sample with HCG at 300IU/L which should have been a clear positive result. This is the second poor result in a row for Children's ward. | CA/PA = EQA/474 | Being investigate 06/05/2021
Fully approved 3/6/21
Web86141 | 2021-04-22 | Urine pregnancy EQA distribution W124: ED had a poor performance for sample 3 where they had reported a negative result when the sample was a spiked urine sample with HCG level of 300 IU/L which should have been a clear positive. | CA/PA = EQA/475 | No investigation 6/5/21 – emailed on 04/05/21 for advice
Fully approved 7/5/21
web86391 | 2021-05-04 | The ED blood gas analyser was in a very dirty state on Friday there was blood over the surface, screen, floor, cupboard and sample port area. People were using it in this state and making no attempt to clean it. It was very difficult to use the analyser without standing in the blood or touching the blood on the analyser. | CA/PA = INC/HS/112 | No investigation 6/5/21
Being reviewed 3/6/21
Finally approved 16/6/21 NH
web86645 | 2021-05-13 | Cardiac ward (glucose) we were unable to find the meter KAAV315-A0125 | CA/PA = EQA/476 | No investigation 3/6/21
|  |  | Ward contacted me in response to non-return letter. Found the meter in day case and have run the sample. I have moved the meter to day case from cardiac ward.(17/05/21 NH) | Investigated 29/6/21 NH.
web86646 | 2021-05-13 | Haygarth ward (glu & ket) KAAV336-A0649. Meter not on ward at time of distribution several suggested wards checked for this meter but unsucsefull. | CA/PA = EQA/476 | Fully approved 18/5/21
web86647 | 2021-05-13 | Pain Clinic (glucose) KAAV329-A0446. Meter being qc'd regularly but not patients. | CA/PA = EQA/476 | Fully approved 19/5/21
web86648 | 2021-05-13 | Pulteney Ward (glucose) KAAW037-A0233. This meter has not been used since October 2020. No record of return to lab for repair. | CA/PA = EQA/476 | Investigation done NH 17/05/2021. Datix not closed yet. Fully approved 25/5/21
web86649 | 2021-05-13 | William Budd day care/oncology OP (glucose) KAAV329-A0122. Meter being used and QC'd regularly | CA/PA = EQA/476 | No investigation 3/6/21
2021-06-29
Web86651 | 2021-05-13 | Main theatre recovery (glucose) KAAX111-A0073. Meter being used regularly, last 2 patients were on millennium as ophthalmology and urology (13/05/2021) | CA/PA = EQA/476 | No investigation 3/6/21
2021-06-29
web86652 | 2021-05-13 | ED (glucose) KAAV329-A0249. This meter is being used regularly in ED. | CA/PA = EQA/476 | No investigation 3/6/21 – changed investigator to Ms F minors
| 2021-06-01 | NICU datix regarding the down time of the BGA analyser on the weekend of the 30/05/2021 – NH did investigation no apparent downtime 2 failed cals one occasion of S2 running out but ward dealt with as would expect. |  | Investigation complete 03/06/21 NH. – finally approved 17/6/21
web87118
web87112 | 2021-06-01 | BBC datix regarding the down time of their BGA from SD wash bypass 2 errors over the weekend 29th – 31st of May. |  | Investigation complete 03/6/21 NH – I approved this as nothing else was added to the investigation 29/06/2021 NH
| 2021-06-02 | Patient attend post operation to his leg. Pt developed swelling to calf. Risk of DVT. D-Dimer test in the department via cassette test “finger prick test” result NAD. Pt returned to Orthopaedic a day later and found to have D-Dimer of 5,000 via venous blood test in laboratory. Therefore pt has a DVT in his affected leg – a life threatening issue if left untreated. | Urgent treatment center – raised by James Brewster (bank nurse) | Still not investigated 29/06/2021 NH
web87183
| 2021-06-06 | Blood gas machine on NICU not working for 2nd consecutive night. | No name given of who raised the datix. | Fully approved 11/06/2021 NH
web87293 |  | Three very sick babies on NICU on respiratory support requiring numerous blood gases. Gas machines on BBC and MAU lactate not working.
Sick patients are left so that nurse can walk long distances to other areas in search of a working machine to process bloods.
web87852 | 2021-06-25 | External quality assurance samples were taken to Children’s ward on 21/05/2021 and were analysed by ward staff on the 1st June. The results were returned to the POCT department and entered onto the WEQAS website for statistical analysis. | CA/PA EQA/485 | Web85743 – not investigated 28/07/2021
|  | The performance was reported on the 10/06/2021 when the pO2 performance was noted to be poor. This has been the case for each monthly distribution this year and two datixs have already been raised concerning this performance (web85736 & web85743). Emails were sent in March and more recently in June with concerns over how these samples are being handled. |  | Web87852 – not investigated 28/7/21
web88890 | 3.8.21 | Children’s ward FPP EQA. | EQA/490
External quality assurance samples were sent to the ward around the 14 July 2021 for the FPP glucose and ketone meters KAAU150-D0400, KAAW068-A0077. The samples should have been analysed and the results returned back to POCT. Unfortunately the samples were not run and even after a couple of phone calls the meters were still not tested.
web88891 | 3.8.21 | Cardiac ward FPP EQA. | EQA/490
External quality assurance samples were sent to the ward around the 14 July 2021 for the FPP glucose meter KAAV313-A0514. The sample should have been analysed on all the meters on the ward and the results returned back to POCT. Unfortunately the samples were not run on this meter despite it being regularly used on cardiac ward.
web88892 | 3.8.21 | Haygarth ward FPP EQA | EQA/490
External quality assurance samples were sent to the ward around the 14 July 2021 for the FPP glucose and ketone meter KAAV336-A0646. The sample should have been analysed on all the meters on the ward and the results returned back to POCT. Unfortunately the samples were not run on this meter and a note to say the meter is lost was made.
web88894 | 3.8.21 | MAU FPP EQA | EQA/490
External quality assurance samples were sent to the ward around the 14 July 2021 for the FPP glucose meter KAAU013-D2164. The sample should have been analysed on all the meters on the ward and the results returned back to POCT. Unfortunately the samples were not run on this meter.
web88895 | 3.8.21 | William Budd day case FPP EQA | EQA/490
External quality assurance samples were sent to the ward around the 14 July 2021 for the FPP glucose meters KAAV329-A0122, KAAX103-A0220. The samples should have been analysed and the results returned back to POCT. Unfortunately the samples were not run and even after a couple of phone calls the meters were still not tested.
web88896 | 3.8.21 | ED FPP EQA | EQA/490
External quality assurance samples were sent to the ward around the 14 July 2021 for the FPP glucose meter KAAV202-A0220. The sample should have been analysed on all the meters on the ward and the results returned back to POCT. Unfortunately the samples were not run on this meter despite it being regularly used on ED.
web89789 | 7.9.21 | Children wards Urine pregnancy EQA for July (UH0721) had one sample reported as negative when it was from a 17 week pregnant urine donor. | EQA/492
|  | ED urine pregnancy EQA for July (UH0721) was not returned and the message samples lost on unit given over the phone. | EQA/494
web89823
Web89821 | 9.9.21 | Theatre 11 returned a poor results for sample 1 hemocue EQA 2108HB | EQA/493 | No investigation 13/10/21
web89822 | 9.9.21 | Main PACU returned a poor result for sample 2 hemocue EQA 2108HB | EQA/493 | No investigation 13/10/21
Web90852 | 2021-10-13 | Theatre 11 (SN1512013016) | EQA/496 – closed 26/10/21 | I added to investigation 27/10/21 – re repeat results
Sample 1 = 169, target = 175.7, DI = -2.10
Sample 2 = 73, target = 77.89, DI = -3.50
web90853 | 2021-10-13 | PAW (1509013033) | EQA/496– closed 26/10/21 | I added to investigation 27/10/21 – re repeat results
Sample 1 = 166, target = 175.70, DI = -3.07
Sample 2 = 75, target 77.89, DI = -2.04
Web91031 | 2021-10-19 | WB - Patient required ketone testing, no machine available on ward. Contacted Waterhouse their machine broken. Parry ward is a covid ward, MAU refused to lend machine. | Senior sister on William budd is investigator and manager | I added to investigation 27.10.21
37